| PPB1 For Fibres and Low Mass Tapes ] |  |  |
| --- | --- | --- |
| Abstract |  |  |
| This document describes the patch panel PPB1 for the SCT fibres and low mass tapes. |  |  |
| Prepared by : | Checked by : | Approved by : |
| V. Cindro, A. Nichols and T. Weidberg |  |  |
| Distribution List |  |  |
| Table of Contents |  |  |

1	Scope of the document	3

2	Introduction	3

3.PPB1

3.1.Description..................................................................................................................................................................	4

3.2 Physical dimensions 	4

3.3 Electrical layout of PP1.......................................................................................................................................	5

3.4 Pads for low mass tape soldering  AND PIN	6

DEFINITION   On   FCI and JST connectors	6

4	References	7

# Scope of the document

This document describes the SCT patch panel PPB1 for the ribbon fibre and low mass tapes.

# Introduction

The first patch panel connection (PPB1) is situated on the LAr cryostat wall (1150mm radius).  Each patch panel assembly caters for two harnesses (twelve tapes).  The aluminum LMTs of the barrel harness will be thermode soldered to PCBs, while the conventional cables will be attached via connectors. The whole assembly is enclosed by an aluminium cover which gives mechanical protection as well as acting as a continuation of the electrical screen.

The fibre ribbons are simply connected at PPB1 using MT12 connectors. There is insufficient space for a bulkhead mounted connector housing, so the connectors are retained by a non-magnetic spring clip[Error: Reference source not found]. The bunch of connections occupies approximately the same Z region as the power tape connections and is held to the cryostat wall by a simple cover/retaining plate. This also fully encloses the area to comply with laser safety guidelines.

# Figure1:View of PPB1, tapes and cables

## Figure 2: One and two PPB1  mounted on the cryostat wall

## 3. PPB1

3.1 DESCRIPTION

PPB1 accommodates the aluminum low mass tapes  on the module side and conventional cables with connectors on the power supply side. One conventional cable accomodates two connectors (4 pin JST for power lines and 20 pin FCI connector for other lines).  The design of PPB1  PCB has following features:

- it provides soldering pads for thermode soldering of the low mass tapes to the printed circuit. Pad pattern was designed after thermode soldering tests done at Rutherford

- it accepts six double layer aluminum low mass tapes, which means that  it serves six modules in one harness

- it provides filtering of the noise on the electrical lines

- the  conventional cable shield  foil will be in electrical contact with a digital ground layer of the PPB1.

- it is designed in eight layers, with three  layers  having a large surface with a shield

- mounting screws are electrically connected to the shield plane

- copper conductor thickness is 35 microns

- the design obeys 2.5 mm distance  (required by IPC-2221 Generic Standard on Printed Board Design) between the HV and other lines on the PCB surface

- several pins will be removed from FCI  connectors  to conform with the HV rule

3.2 PHYSICAL DIMENSIONS

width 38  0.2 mm

length 268   0.3mm

PCB thickness  2.2  0.2 mm

3.3 ELECTRICAL LAYOUT OF  PPB1

The  electrical design of the board is based on the studies made in the system test and is called PPB1-V6Z. The schematic is shown in Figure 3.  It provides filtering of HF noise. PPB1-V6Z is an evolution of PP1N-V3  and PP1W-V3 assembly used in the system test with the following changes:

- integration of two boards into one board with 8 layers

- connectors  (4 pin FCI and 20 pin JST) will be used to connect type II  cables (Raydex AWG 20).

Figure 3: Schematic for  PPB1

Filtering of noise is provided  with capacitors. Main features are:

- 4.7F capacitors from  Vcc, Vdd to  their return lines (Murata GRM31CR61A475KA01K – size 1206, 10V)

- 15 nF from  HV to  its return line (Murata GRM31CR72J153KW03L – size 1206,  630V)

- 0.1F  from  sense, their returns, reset lines  to the  shield layer (Murata  GRM21BR71E104KA01K – size 0805, 25V)

- 1 F from  Temp1, Temp2, Vcsela, Vcselb, Select, Pin bias, HVgnd and  Agnd

to the  shield plane (Murata Murata  GRM21BR71C105KA01K – size 0805, 16V)

-     zero ohm resistor  between Dgnd  and shield

## 3.4 Pads for low mass tape soldering  AND PIN

## DEFINITION   On   FCI and JST  CONNECTOR

The pads for the layout of the thermode soldering  of the LMTs to PPB1 are shown Figure 4 in below.

Figure 4 Layout of pads on PP1B

Figure 5: Pin definition on FCI and JST connectors

# References

Low Mass Tapes, EDMS ATLS-IS-AT-0010

2 MT Connectors, Spring Clips and Infineon Connectors, EDMS ATL-IS-ES-0053.