ICATPP 2009 - 11th International Conference on Advanced Technology and Particle Physics
5th - 9th October 2009, Villa Olmo, Como, Italy
The ATLAS
Beam Condition and
Beam Loss Monitors
Irena Dolenc
(Jožef Stefan Institute, Ljubljana)
on behalf of the ATLAS BCM group
[Speaker Notes: Theme created by Sakari Koivunen and Henrik Omma Released under the LGPL license.]

---
ATLAS BCM group
JSI Ljubljana
V. Cindro
I. Dolenc
A. Gorišek
G. Kramberger
B. Maček
I. Mandić
E. Margan
M. Mikuž
M. Zavrtanik
CERN
D. Dobos,
H. Pernegger,
C. Spartharakis
E. Stanecka,
P. Weilhammer
Ohio State University
H. Kagan
S. Smith
J. Moss
Univ. of Applied Science,
Wiener Neustadt
E. Griessmayer
H. Frais-Kölbl
M. Niegl
Univ. of Toronto
C. Cadabeschi
E. Jankowski
D. Tardif
W. Trischuk
ATLAS BCM web page: https://twiki/cern.ch/twiki/bin/view/Atlas/BcmWiki

---
Motivation
LHC equipment failure: resulting beam losses potentially dangerous to the ATLAS Inner Detector (ID)
Experience show that beam accidents can happen
Total energy stored in one LHC beam (2808 bunches with 1011 protons at 7TeV → 350MJ) more than 100-times higher than in previous accelerators like Tevatron or HERA
Time constants of magnets: shortest ~ few LHC turns (order of ms) → if beam losses detected early enough, the beam can be dumped in time to prevent beam accident (beam is dumped within 3 turns ~ 270μs)
Tevatron Beam accident:
caused by Roman Pot reinserting itself in the beam after it had been issued the retract commands
Secondary collimator (stainless steel) for halo removal
20-30 Tevatron turns

---
ATALS BCM and BLM goal
Protection
Passive:
ATLAS and CMS have TAS collimators: protects Inner Detector from direct beam incidences
TAS (Target Absorber Secondaries) collimators: at z=±18m which protect the Inner Triplet of quadrupoles from secondaries produced in p-p collisions
Active:
Machine Beam Loss Monitors, Beam Positions Monitors... can fire beam abort
ATLAS Beam Condition Monitor (BCM) and Beam Loss Monitor (BLM)
Designed for fast detection of early signs of beam instabilities (due to incorrect magnet settings, magnet trips, failures...)
Issue a beam abort in case of beam failures
BCM
Particle counter
Will additionally provide a coarse relative luminosity measurement (complementary information to LUCID-ATLAS main luminosity monitor)
BLM
Induced current measurement

---
ATLAS BCM: principle of operation
Time of flight measurement to distinguish between collisions and
background events (beam gas, halo, TAS scraping)
measurement every proton bunch crossing (BC - 25 ns)
place 2 detector stations at z=1.9m:
particles from collisions reach both stations at the same time (6.25ns after collisions themselves) every BC
particles from background interactions
reach the nearest station 12.5ns (½ BC) before particles from collisions (6.25ns before collisions themselves)
reach the furthest station at the same time as particles from collisions
use “out of time” hits to identify the background events
use “in time” hits to monitor luminosity
Requirements:
fast and radiation hard detector & electronics:
rise time ~1ns
pulse width ~3ns
baseline restoration ~10ns
ionization dose ~0.5 MGy, 1015particles/cm2 in 10 years
MIP sensitivity

---
BCM: Realization
4 BCM detector modules on each side of the Interaction Point
Mounted at z=183.8cm and =0o, 90o, 180o, 270o, sensors at r=5.5cm (≈4.2)
Modules tilted by 45o towards the beam pipe
ATLAS Inner
Detector
TRT
ATLAS
SCT
Pixel
Courtesy of
T. Kittelmann
BCM detector
modules

---
BCM: Detector Modules Installed
BCM modules were installed on Beam Pipe Support Structure (BPSS) in November 2006 and lowered into ATLAS pit in June 2007
Beam
Pipe
BCM
module
Pixel
Detector

---
BCM Detector modules
Poly-crystalline CVD diamond sensors
developed by RD42 and Element Six Ltd.
radiation hard: shown to withstand 1015p/cm2
low leakage current  no cooling required
operated at high drift field 2V/μm → fast and short signals
Double – decker assembly
2 back-to-back sensors each with
Thickness 500μm,
CCD @1V/cm ~220μm
Size: 10×10 mm2
Contact size: 8×8 mm2
Operated at 2V/μm (1000V)
Double signal compared to assembly with
one sensor, but noise not measured to be
two times higher
For 45o particle incidence signal increase
by factor √2  modules installed at 45o
to the beam pipe
13.4 cm

---
BCM Detector modules
Front end electronics
2 stage amplifier:
1st stage: Agilent MGA-62653, 500MHz (22db)
2st stage: Mini Circuit GALI-52, 1GHz (20dB)
Measurements showed (confirmed with simulation):
Limiting BWL to 200MHz improved S/N by 1.3
rise time worse by 70% and FWHM by 60%, but still fit to requirements
 4th order 200MHz filter integrated on digitisation board before digitisation (NINO chip)
13.4 cm

---
BCM: Readout chain
Analogue signals from BCM detector modules routed behind calorimeter (lower radiation levels) where they are “digitised” by custom made board based on NINO chip
Each module connected to separate NINO electronics board
Counting room
USA15
PP2 (~10Gy)
BPSS
~0.5MGy

---
BCM: “Digitisation” electronics board
NINO chip
Developed for ALICE ToF (F. Anghinolfi et al.)
Radiation tolerant, fabricated in 0.25μm IBM process
Rise time < 1ns, jitter <25ps, Min. detection threshold 10fC
Time-over-threshold amplifier-discriminator (adjustable threshold) chip
Width of LVDS output signal depends on input charge
Before input to NINO chip: signal charge split in two channels in ratio of 1:11 to increase the dynamic range → high and low gain channels
Low gain channels need ~10 times more MIPs traversing the sensor simultaneously than high gain channels to show the signal
NINO
NINO
output
input signal
(BCM)

---
BCM: Readout chain
Optical signals transmitted through 70m of optical fibres to USA15 counting room
Optical signals transformed to PECL (2 optical receiver boards) and connected to 2 data processing units based on Xilinx Vitrex-4 FPGA
Processing units connected to
ATLAS Central Trigger Processor (CTP): BCM provides 9 bits
ATLAS Data acquisition (TDAQ): on trigger signal (LVL1A) from CTP, BCM data from 31 bunch crossings is formatted and sent of ATLAS DAQ readout chain
ATLAS Detector Control System (DCS): monitoring rates,...
Detector Safety Systems (DSS), Beam abort system (through CIBU-Control Interlocks Beam User)
Counting room
USA15
PP2 (~10Gy)
BPSS
~0.5MGy

---
ATLAS BCM: Cosmic data
November 2008: combined ATLAS Inner Detector cosmic data taking
Two different triggers:
Resistive Plate Chambers (RPC) of Muon system
Fast-OR mechanism of Transition Radiation Tracker (TRT)
On each trigger from CTP: BCM sends data (signal rising edge and width) from 31 BCs to ATLAS DAQ chain
Timing distribution of RPC and TRT triggered BCM hits:
Superimposed is a fit to Gaussian signal and a random background.
Gaussian peak wider with RPC trigger: better jitter of TRT trigger
Probability for noise hit in BC: ~10-7
RPC: about 10M triggers needed for 9 true BCM hit
TRT: about 1M triggers needed for 9 true BCM hit

---
ATLAS BCM: Cosmic data
June 2009: ATLAS cosmic data taking
IDCosmic trigger stream: selects events with at least 1 track reconstructed in Inner Detector at trigger LVL2
RNDM trigger stream: random trigger
Timing plots:
RNDM stream: flat as expected
IDCosmic stream: Gaussian width ~ as for TRT stream in November cosmic run
Probability for noise hit per 1 BC: 8-7 times higher then in November 2008 run, due to higher thresholds in 2008
Signal: 1M IDCosmic triggers give 6 true BCM hits

---
ATLAS BCM: Cosmic data
Timing distribution of BCM hits over channels:
Almost no hits on C side (attributed to unequal size of 2 main ATLAS shafts)
TRT
From: M. Bianco,”ATLAS RPC commissioning status and cosmic ray test results”, PIC2008
IDCosmic
RPC

---
ATLAS BLM
ATLAS Beam Loss Monitor – ATLAS BLM
Independent system
Recently added as a backup for ATLAS BCM
Goal: only protection → much simpler than BCM
Provide complementary information to BCM, but can be used as standalone in case of problems with BCM
Readout electronics based on the LHC BLM system
Major change: ATLAS BLM uses diamond sensors instead of ionisation chambers (LHC BLM)
Other: different LV powering scheme, firmware changes in the readout cards,...
Sensors:
one pCVD 8×8mm2 diamond, 500μm
thick. metallization 7×7mm2
operated at +500V
current @500V typically <1-2pA

---
ATLAS BLM: Detector modules
12 detector modules installed in May 2008
6 on each side of IP, installed on Inner Detector End Plate
Close to IP: z=±3450mm, r=65mm
Inner Detector End Plate
ATLAS Inner
Detector
Courtesy of
T. Kittelmann

---
ATLAS BLM: Readout
Measurement of radiation induced current in sensors, integrated over predefined time constants ranging from 40μs to 84s.
Digitization of induced current: converted to frequency by radiation tolerant BLMCFC (“CFC card”, behind calorimeter)
Digitised information transmitted through optical fibers to USA15 counting room, recorded by BLMTC (“threshold card”) based on Altera FPGA:
BLMTC inserted in VME crate
Through Single Board Computer, VME bus used recording the post mortem buffer and for monitoring readings (sent to ATLAS Detector Control System)
Connected to LCH beam abort: beam abort issued if readings from 2 modules on one side exceed a predefine threshold
BLMCFCs
BLMTC

---
Summary
ATLAS BCM will monitor beam conditions
close to IP using TOF measurement
Goal:
Protection
Additionally, relative luminosity measurement
pCVD diamonds as sensor material
2 diamonds in a back-to-back configuration
at 45o towards the beam
First experience with the system obtained
in the last year and a half
ATLAS BLM:
Redundant system for safety purposes
Induced current measurement
pCVD diamond sensors
Looking forward to using them in the real LHC environment

---
Backup slides

---
Beam Loss Scenarios
Single-turn losses
Likely to occur during injection or beam dump processes
ATLAS can be rated the 'safest' of all LHC interaction points (far away from injection, dump)
Pilot bunch (single bunch of low intensity, 5×109p@450GeV; 360J) will be used to check the magnet settings
Multi-turn losses
Beam degradation due to equipment failure, magnet trips, wrong magnet settings...
Time constants of magnets: shortest ~ few LHC turns (order of ms) → can abort the beam if detected early (beam is dumped within 3 turns ~ 270μs)
Simulation of beam orbits with wrong magnet settings (D. Bocian) exhibit scenarios with pilot beam scrapping the beam pipe or TAS collimator (most likely scenario), no direct hitting of Inner Detector (ID)
TAS (Target Absorber Secondaries) collimators: at z=±18m which protect the Inner Triplet of quadrupoles from secondaries produced in p-p collisions
60 m
30 cm

---
QA of BCM Modules
Qualification tests with final modules to select 8 for installation
Raw sensor characterization:
I/V, CCD
Module performance
All modules subjected to thermo-mechanical test
infant mortality test (12h @80o)
accelerated ageing for one of the module
(14h @120o 10 years at 20o)
thermal cycling (10 cycles from -25o to 45o)
Module performance checked before and after
these tests with 90Sr setup
no change in S/N observed
for normal particle incidence: typical S/N7–7.5
(trigg.)

---
BCM Detector modules
Front end electronics
2 stage amplifier:
1st stage: Agilent MGA-62653, 500MHz (22db)
2st stage: Mini Circuit GALI-52, 1GHz (20dB)
Limiting BWL to 200MHz improved S/N by 1.3 (and rise time worse by 70% and FWHM by 60%, but still fit to requirements)
 4th order 200MHz filter integrated before digitization (on NINO board)
200MHz BWL
1GHz BWL
Amplitude dist.
Noise dist.
Lab. Tests with
MIPs at 90o
incidence
Off-line analysis of waveforms recorded at full BWL:
optimum S/N with 1st order filter with cut-off frequency 200-400 MHz
13.4 cm

---
BCM: Analogue signals
Signals recorded with 200MHz BWL at the readout
Typical S/N with MIP at 90o incidence in the lab. tests: 7-7.5
Most probable amplitude produced by MIP at 90oin the lab tests: 2.2 - 2.7mV
Mean rise time 1.4ns
Mean FWHM 2.9ns
Timing resolution better than 400ps (thresholds: 0.1-2 MP amplitude)
200MHz BWL
1GHz BWL

---
BCM: luminosity
BCM will contribute to luminosity monitoring with
Monitor instantaneous luminosity
Vertex position monitoring
Determine dead time
Beam separation scans
First algorithms will be based on non-empty event counting
Monitoring of luminosity per BCID
Providing instantaneous luminosity at Hz rate
Monte-Carlo simulations under-way to provide initial calibration used before first beam-separation scans, and understanding systematic

---
BCM: Processing units & signal analysis
Input signal sampled with 2.56 GHz:
64 samples of 390ps width for each proton bunch crossing (BC-25ns)
Raw data stored in DDR2 circular buffer (for >more than 3×106 BCs → 1000 last LHC turns)
For each BC:
Signal rising edge and width of at most first 2 signals are reconstructed and stored in DDR cyclic buffer (more than 2×1061BCs → 800 LHC turns)
On trigger signal (LVL1A) from CTP: BCM data from 31 Bunch crossings is formatted and sent of ATLAS DAQ readout chain
“In-time” and “out-of-time” coincidences, high multiplicity for low and high gain channels
9 trigger bits to CTP
Beam abort (if beam conditions have reached unacceptable level)
Alarm and warning signals in less sent to DSS, DCS

---
BCM: beam abort
Beam abort algorithms:
For initial phase: 4 Low threshold channels (horizontal modules) + 3 high threshold channels (vertical modules)
Monitor and use later:
X/Y (3+4 condition reached in consecutive X out of Y BCs)
Leaky bucket (X/Y with forgetting factor)

---
BLM: induced current
Currents in BLM modules:
Current from 14TeV p-p collisions:
~15nA/module (at design luminosity, 1034/cm-1s-1)
Estimated form the Total Ionisation Dose plots
Scales with luminosity
Accidents (simulation)
7 TeV p on TAS gives ~1 MIP/BLM module
1 MIP in diamond generates ~1 fC of charge
25pA of current “spike” for single occurrence (1 bunch, possible for pilot bunch)
40nA of continuous loss every BC (much more likely for full LHC bunch structure)
Diamond dark currents in magnetic field O(10pA)
Initially threshold for 40μs reading: set to 50nA

---
BCM: performance up to FPGA input
Estimated performance of BCM system in ATLAS (up to FPGA input) for high gain channels:
Median S/N (for MIPs at 45o incidence)  9
Noise RMS σ  64mV
Median signal  570mV
1MIP efficiency versus noise occupancy (noise rate scaled to 25ns interval – bunch crossing in ATLAS)
efficiency 0.95 – 0.99 for occupancies 10-3–10-4
The exact level of fake rate depends on what kind of logical combination of signals will be used in ATLAS

---
BCM: performance
Noise rate measured in ATLAS pit:
extracted noise RMS σ 51-68mV (different det. modules and NINO boards) agrees with estimation (64mV)