H6 Test beam
(29. okt – 12. nov. 2007)
Irena Dolenc, Andrej Gorišek

---
Setup
ADC ch0: F401 (BCM2) ->irradiated with
  - 2.3 1014p/cm2
ADC ch1: F409 (BCM3)
ADC ch2: F406 (BCM4)
ADC ch3: F413 (trigger,
NIM signal, attenuated by
factor 0.25 ie. 12db)
trigger: coincidence of scintillators 1,2,3,4
Oznake:
NINO1:	brez obeh uporov, brez filtra
NINO2:	Helmutova verzija
NINO3:	Erikova verzija
NINO1mod:	dodan filter (9. nov. 2007)

---
NINO efficiency & noise rate:
Oznake:
NINO1
NINO2
NINO3
NINO1mod
modri markerji: F401
rdeci markerji: F409
zeleni markerji: F406
polni markerji: NINO ch1
prazni markerji: NINO ch3

---
NINO1:
σ from the fit: log(Noise Rate)=A+
2σ2
Thr2
Module	σ [mV] Thr @ 0.5 eff [mV]
ch1	ch3 ch1 ch3
F401	46.2, 45.4	47.8 380 380
F406	/ 49.8 350 370
F409	51.6 53.3 340 370

---
NINO2:
Module	σ [mV] Thr @ 0.5 eff [mV]
    - ch1	ch3 ch1 ch3
F401	31.9, 31.5 31.5, 32.8 250 250
F406	33.5 33.1 265 250
F409	38.3 38.1 210, 230 210, 230

---
NINO3:
Module	σ [mV] Thr @ 0.5 eff [mV]
ch1	ch3 ch1 ch3
F401	9.3 9.9 105, 105 110
F406	7.3 8.7 130, 110 135, 115
F409	/ / 85 93, 95

---
NINO1 mod.:
Module	σ [mV] Thr @ 0.5 eff [mV]
ch1	ch3 ch1 ch3
F401	27.9 26.0 225 235 ?
F406	31.7, 32.8 32.4, 32.9 245 260
F409	36.7 31.2 200 ? 225

---
F401:

---
F409:

---
F406:

---
NINO SNR:
Module NINO1 NINO2 NINO3 NINO1 mod.
ch1 ch3 ch1 ch3 ch1 ch3 ch1 ch3
F401	8.3 7.9 7.9 7.8 11.3 10.6 8.1 9.0 ?
F406	/ 7.4 7.9 7.6 16.4 14.4 7.6 8.0
F409	6.6 6.9 5.7 5.8 / / 5.4 ? 7.2

---
Efficiency Vs. Threshold: analogue signal @+1000V
F409
F401
TB julija 2007 (H8), F401 neobsevan
TB avgusta 2007 (H6), F401 obsevan z 1.02*1014 p/cm2
zadnji TB oktobra 2007 (H6), F401 obsevan z 2.37*1014 p/cm2 (skupna fluenca)
F406

---
Signal & Noise: analogue signal @+1000V
signal
noise
F401
signal sampled in 2.5 ns interval around average signal
crna: TB julija
rdeca: TB v agustu
modra: zadnji TB
F409
F406

---
Rezultati iz TB v Avgustu:
NINO2 signali: noise rate
black markers: F409
red markers: F406
blue markers: F401
green markers: nothing connected
open/closed markers: different NINO channels (ch1,ch3)
σ from the fit: log(Noise Rate)=A+
Module	σ [mV]
F401	32
F406	31
F409	43
/	12(ch1), 15(ch2)
2σ2
Thr2

---
Rezultati iz TB v avgustu
NINO2 signals: efficiency
searching for signal in 50ns time interval
black markers: F409
red markers: F406
open/closed markers: NINO inputs swapped
???
persist. plot for run with 200mV threshold:
F409
NINO ch1
F401
analogue signal
F406
median efficiency at ~ 335mV and σ from noise rate fit give SNR~10
from QA: SNR~10 if normalized to 45o
trigger
F406
NINO ch3
F409
median efficiency at ~ 195mV and σ from noise rate fit give SNR~4.5
from QA: no data for pos. polarity

---
Rezultati iz TB v avgustu:
NINO2 Efficiency Vs. Occupancy for F406
efficiency calculated from ADC data
occupancy (in 25 ns time interval) calculated from fit to Noise rate measurement for a given efficiency data point
NINO2 Thr [mV]	Eff.	Occ.
  - 180.4	0.99	2.9·10-8
  - 200.4	0.97	6.7·10-9
  - 225	0.95	3.7·10-12
  - 250.2	0.90	9.9·10-15
  - 300.4	0.70	1.1·10-20