Vertex 2008 - 17th International Workshop on Vertex Detectors
27th July - 1st August 2008, Üto, Stockholm, Sweden
ATLAS Diamond
Beam Conditions Monitor
Irena Dolenc
on behalf of the ATLAS BCM Collaboration
Presented in Irena's absence by Harris Kagan
[Speaker Notes: Theme created by Sakari Koivunen and Henrik Omma Released under the LGPL license.]

---
ATLAS BCM collaboration members
JSI Ljubljana
V. Cindro
I. Dolenc
A. Gorišek
G. Kramberger
B. Maček
I. Mandić
E. Margan
M. Mikuž
M. Zavrtanik
CERN
D. Dobos,
H. Pernegger,
P. Weilhammer
Ohio State University
H. Kagan
S. Smith
Univ. of Applied Science,
Wiener Neustadt
E. Griessmayer
H. Frais-Kölbl
M. Niegl
Univ. of Toronto
C. Cadabeschi
D. Tardif
W. Trischuk
ATLAS BCM web page: https://twiki/cern.ch/twiki/bin/view/Atlas/BcmWiki

---
Motivation
The goal of BCM system inside the ATLAS Inner Detector:
Monitor beam conditions and distinguish each bunch crossing between normal collisions and background events during normal running
measure background rate (beam halo, beam gas) close to the vertex
measure collision rate and provide (bunch by bunch) relative luminosity information (additional measurement to LUCID, ATLAS main luminosity monitor)
Primary goal: protection in case of larger beam losses
fast detection of early signs of beam instabilities
(due to incorrect magnet settings, trips, ...)
Issue warning and alarm signals for equipment
protection
Input to ATLAS Detector Safety System and
LHC Beam Abort

---
Beam Loss Scenarios
Multi-turn losses
beam degradation (equipment failure, wrong magnet settings...)
time constants of magnets ~ms → can abort the beam if detected early
Single-turn losses of 360MJ/beam
likely to occur during injection or beam dump processes
IR1 (ATLAS) can be rated the 'safest' of all interaction points (far away from injection, dump)
pilot bunch (single bunch of low intensity, 5×109p@450GeV; 360J) will be used to check the magnet settings
simulations of beam orbits with wrong
magnet settings (D. Bocian) exhibit
scenarios with beam scrapping TAS Cu
collimator
60 m
30 cm

---
ATLAS BCM principle of operation
Time of flight measurement to distinguish between collisions and
downstream background events (beam gas, halo, TAS scraping)
measurement every proton bunch crossing (25 ns)
place 2 detector stations at z=1.9m:
secondary particles from collisions reach both stations at the same time (6 ns after collisions)
secondary particles from upstream background interactions reach nearest station 12.5ns before secondary particles from collisions (6 ns before collisions)
use “out of time” hits to identify the background events
use “in time” hits to monitor luminosity
Requirements:
fast and radiation hard detector & electronics:
rise time ~1ns
pulse width ~3ns
baseline restoration ~10ns
ionization dose ~0.5 MGy, 1015particles/cm2 in 10 years
MIP sensitivity

---
Realization
4 BCM detector modules on each side of the Pixel detector
Mounted on Beam Pipe Support Structure at z=183.8cm, sensors at r=5.5cm (≈4.2)
BCM detector
modules

---
BCM Detector Modules Installed
BCM modules were installed on Beam Pipe Support Structure in November 2006 and lowered into ATLAS pit in June 2007
Beam
Pipe
BCM
module
Pixel
Detector

---
BCM System - Schematics
Counting room
USA15
PP2
~10Gy
BPSS
~0.5MGy

---
BCM Detector modules
Poly-crystalline CVD diamond sensors
developed by RD42 and Element Six Ltd.
radiation hard: shown to withstand 1015p/cm2
low leakage current  no cooling required
fast & short signals – operate at high drift field 2V/μm
Double – decker assembly
2 back-to-back sensors each with
thickness 500μm,
CCD @1V/cm ~220μm
Size: 10×10 mm2
Contact size: 8×8 mm2
Operated at 2V/μm (1000V)
Double signal compared to assembly with
one sensor, but noise not measured to be
two times higher
For 45o particle incidence signal increase
by factor √2  modules installed at 45o
to the beam pipe
13.4 cm

---
BCM Detector modules
Front end electronics
2 stage amplifier:
1st stage: Agilent MGA-62653, 500MHz (22db)
2st stage: Mini Circuit GALI-52, 1GHz (20dB)
Limiting BWL to 200MHz improved S/N by 50% (but 10% worse timing resolution)
 4th order 200MHz filter integrated on NINO board before digitization
Signals recorded with LeCroy LC564A digital oscilloscope with 200MHz BWL:
mean rise time 1.4ns
mean FWHM 2.9ns
13.4 cm

---
NINO board
NINO chip
Developed for ALICE ToF (F. Anghinolfi et al.)
Radiation tolerant
Fabricated in 0.25μm IBM process
Peaking time< 1ns, jitter <25ps
Min. detection threshold 10fC
Time-over-threshold amplifier-discriminator chip
width of LVDS output signal depends on input charge
NINO
output
input signal
(BCM)

---
QA of BCM Modules
Qualification tests with final modules to select 8 for installation
Raw sensor characterization:
I/V, CCD
Module performance
All modules subjected to thermo-mechanical test
infant mortality test (12h @80o)
accelerated ageing for one of the module
(14h @120o 10 years at 20o)
thermal cycling (10 cycles from -25o to 45o)
Module performance checked before and after
these tests with 90Sr setup
no change in S/N observed
for normal particle incidence: typical S/N7–7.5
(trigg.)

---
Test Beam at CERN SPS/PS in 2006:
Setup
Bonn Telescope:
4 Si-strip tracking modules
(4 XY points along pion path)
Trigger scintillators
Pion Beam
T9: p=3.5G eV/c
T11: p=12 GeV/c
H8: p=20,100 GeV/c
BCM detector modules

---
Test Beam at CERN SPS/PS in 2006:
xy efficiency scan with analogue signals
Moving in steps of 0.1mm across the xy surface, calculating the fraction of tracks with analogue signal above chosen threshold in 0.2×0.2 mm2 square at each step:
non-uniformity due to the grain structure of the diamonds in assembly
Thr=0.5 MP amplitude
Thr=0.75 MP amplitude
Thr=1 MP amplitude
Module
F408
Module
F403

---
Test Beam at CERN SPS in 2007
NINO board performance
Efficiency curve with triggering on an incident MIP measured by varying the discriminator threshold on NINO board
median signal at 100-130mV with different modules
final boards: new amplification added (Mini Circuit GALI-52) to make system more manageable
Noise rate measured in the lab:
with spare detector module (F406, typical of those installed in ATLAS)
 RMS noise =17mV, median signal ≈125mV  median S/N≈7.3
Noise Occupancy in bunch crossing
1MIP Efficiency

---
Test Beam at CERN SPS in 2007
Timing resolution after TOT
Calculated time difference between NINO pulses
RMS of time difference distributions bellow 1.1ns for scanned threshold range
 time resolution for one module < 780ps

---
Performance of final NINO boards:
Efficiency, Noise Rate
Final NINO boards: with additional amplification
Noise RMS estimated from noise rate measurement in lab with spare final NINO board and module F406 σ=75mV
Efficiency curve measured with spare NINO board and module F406
90Sr source, trigger on BCM detector module signal (non MIP efficiency curve)
measured curve for final board scaled to curve for 1MIP (scaling factor calculated by comparing test beam and 90Sr measurements obtained for the board tested in test beam):
median efficiency  590mV, median S/N7.9
1MIP efficiency versus noise occupancy (noise rate scaled to 25ns interval – bunch crossing in ATLAS)
efficiency 0.95 – 0.99 for occupancies 10-3–10-4
The exact level of fake rate depends on what kind of logical combination of signals will be used in ATLAS

---
Performance of installed NINO boards
Noise rate measured in ATLAS pit (after installation of NINO boards in ATLAS)

---
BLM - Redundant System
BLM – Beam Loss Monitor
Mostly copied from BLM system installed in BaBar
Sensors:
one pCVD 8×8mm2 diamond, 500μm thick. metallization 7×7mm2
operated at +500V
current @500V typically <1-2pA
12 diamonds (6+6) installed on Inner Detector End Plate at pixel PP1 in May 2008
12m coax to PP2 (radiation tolerant BLM cards)
Digitization of ionization current in diamond over 40μs, converted to frequency at PP2 (LHC CFC cards, radiation tolerant)
Optical transmission to USA15 counting room, recorded by Xilinx FPGA
Use standalone in case of BCM troubles, otherwise complementary information

---
Summary
ATLAS BCM will monitor beam conditions close to
IP using TOF measurement
pCVD diamonds as sensor material
2 diamonds in a back-to-back configuration
at 45o towards the beam
Test beam and lab results indicated operable system
rise time, pulse width, S/N meet criteria
efficiency, noise occupancy acceptable
Status:
BCM and BLM (redundant system) installed
focus shifted towards development of FPGA algorithms